PRS1 Data File Format

Each PR System One data file has a well defined structure containing a sequence of one or more data blocks. Each block begins with a 15 byte header, any additional header data, an 8 bit header checksum, the data stream itself (in a not-so-well-defined structure), followed by a 16bit data integrity checksum.

The cycle repeats for each subsequence block, although only waveform data has been observed to contain multiple blocks.

The file's extension provides an extra indication of the type of data contained:

Extension	Type
.001	Session Summary
.002	Event Data
.004	Time details ???
.005	Waveform Data

The filename component minus the extension gives session sequence number. In most cases it would be better to use the internal values instead, but the are handy for quick lookup.

Standard PRS1 Header Example:
[02] [48 0d] [00] [05] [00] [02] [1f 00 00 00] [15 ae 12 4e]

Idx	Data	Type	Details
00	[02]	8bit integer	DataFormatVersion
01	[48 0d]	16bit integer	BlockLength (in bytes, including header & checksum values)
03	[00]	8bit integer	FileType (Defining additional header information)
04	[05]	8bit integer	Family
05	[00]	8bit integer	Family Version
06	[02]	8bit integer	File Extension
07	[1f 00 00 00]	32bit integer	Session Sequence Number
0b	[15 ae 12 4e]	32bit integer	TimeStamp (UNIX Epoch)

Known FileType values
00 indicates no additional header data is needed. The standard header is just followed by the 8bit checksum at position 0x0f, which is additive sum of all header values up to this byte.
01 indicates waveform data storage.

Waveform Data Header (.005 files)

Idx	Data	Type	Details
0x0f	[xx xx]	16bit integer	Number of seconds
0x11	[01]	8bit integer	Unknown
0x12	[xx xx]	16bit integer	Number of signals (nsig)

Followed by a list containing (nsig * 24bits) of data describing how each waveform data is stored:
0x13 [xx xx]	16bit integer	Sample interleave
0x15	[xx]	8bit integer	Sample format
.. repeats (nsig times)

Important: For some reason, this list is stored in reverse order.

Followed by the 8bit sum of all header information, stored at 0x16+(nsig*3)

Calculating just the size of a blocks data area
DataSize = BlockLength – HeaderLength – Extra Header stuff – 1 byte header checksum – 2 byte block checksum

For summary & event data, this translates to DataSize = BlockLength-18
For waveform data, this translates to DataSize = BlockLength-18-(nsig*3)

Known Sample Formats
00	8bit Signed
01	8bit Unsigned

Waveform Data Structure
If only one waveform is present (indicated by number of signals == 1), this simply contains DataSize/SampleSize samples for this block in order.

If more than one waveform is present, the data is multiplexed together in small groups of samples, the size of which is determined by the interleave value and sample type stored in the header.

Only 8 bit samples have been observed, but I feel it's safe to assume the interleave counts samples and not bytes. Different sample types can be interleaved together, and may therefore be different sizes.

Note: Interleaved data is stored in the opposite order to the list data defined in the waveform header component.

Example, Two multiplexed waveforms, one with an interleave of 5, sample type 0, and the other has an interleave of 5, sample type 1

Byte	Value	Waveform
0x00	44	1	(Starting from the first byte in the data block)
0x01	45	1
0x02	46	1
0x03	48	1
0x04	50	1
0x05	00	2
0x06	01	2
0x07	02	2
0x08	03	2
0x09	04	2
0x0a	55	1
0x0b	55	1
0x0c	55	1
0x0d	55	1
0x0e	55	1
0x0f	05	2
... and so on..

Interleaving repeats for the remainder of the DataSize

Event Data Format (.002 files)

Event data is stored after the standard header in time-delta format.

Each event is represented by an 8bit code, and generally followed by 16bits worth of delta offset values, and zero or more data fields of varying byte lengths, determined by the event code..
(I say generally because occasionally events have a short form based on the content.)

The schema describing this information is not stored with the event data, so understanding of the DataFormat, Family & Family Version fields is needed to interpret this file correctly.

System One ASV Data (family 5) for instance, uses different codes than a System One Remstar Auto (family 0)
ASV does not contain RERA (or at least according to Philips Respironics Encore software, you'd think their top of the line machine would have this if it wasn't a bullshite figure to begin with.)

Family 0 Codes
Code	Name	Delta	Data Fields	Notes
0x01	Unknown	[xx xx] No data records
0x02	Pressure	[xx xx] Pressure=[xx]	Divide pressure value by 10
0x03	BIPAP Press.	[xx xx] EPAP=[xx] IPAP=[xx]	Divide pressure values by 10
0x04	PressurePulse	[xx xx] Unknown=[xx]
0x05	RERA	[xx xx] Offset=[xx]
0x06	Obstructive	[xx xx] Offset=[xx]
0x07	ClearAirway	[xx xx] Offset=[xx]
0x0a	Hypopnea	[xx xx] Offset=[xx]
0x0c	FlowLimitation [xx xx] Offset=[xx]
0x0d	Vsnore	[xx xx] No data records
0x0e	Unknown	[xx xx] [xx] [xx] [xx]	Unknown purpose
0x0f	CSR/PB	[xx xx] Duration=[xx xx] Offset=[xx]
0x11	GraphData	[xx xx] Leak=[xx] Snore=[xx]
0x12	Summary?	?????	[xx] [xx] [xx] [xx]

Family 5 Codes (ASV)
Code	Name	Delta	Data Fields	Notes
0x00	Unknown	Variable length..	Freaking weird. Read all the zeros in sequence, and then one more value.
0x02	Pressure	[xx xx] Pressure=[xx]	Divide pressure value by 10
0x04	PressurePulse?	[xx xx] ?? = [xx]	This data looks like pressure pulse.
0x05	Obstructive	[xx xx] Offset=[xx]
0x06	ClearAirway	[xx xx] Offset=[xx]
0x07	Hypopnea	[xx xx] Offset=[xx]
0x09	FlowLimitation [xx xx] Offset=[xx]
0x0b	CSR/PB	[xx xx] Duration=[xx xx] Offset=[xx]	Multiply Duration by 2
0x0d	GraphData	[xx xx] [xx xx xx xx xx xx xx xx xx xx]	These 10 bytes are broken down further into graph data
0x0e	Unknown	[xx xx] [xx]	Unknown Purpose.. RERA?

Delta values are applied before adding the event.
Offset values are subtracted from the running delta (but they don't adjust the running delta, only the individual event)

Example 0x0d GraphData breakdown:

0d [78 00] [50] [50] [51] [1d] [13] [64] [0d] [44] [00] [2e]
      0 1 2 3 4 5 6 7 8 9 10

Idx	Details	Notes
00	The usual 16bit Time-Delta
01	IPAP Value	Divide by 10
02	IPAP Low Range	Divide by 10
03	IPAP High Range	Divide by 10
04	Leak
05	Breaths Per Minute
06	Patient Triggered breaths %
07	Minute Ventilation
08	Tidal Volume / 10.0	Multiply by 10
09	Snore Graph	(>0 = event flag on encore)
10	EPAP Value	Divide by 10

IPAP High – IPAP Low = PS (Pressure Support)?
IPAP – EPAP = Pressure Support?

Summary Data (.001 files)

Summary data is stored after a standard data header, and contain totals, pressure settings, and general machine settings for a recorded session.
Usually they are around 77 bytes, although this varies.

The structure is messy and not well defined, nor is it well understood

A non zero value in the 6th data position inserts 4 bytes directly afterwards..

[ STILL TO DO ]

.004 files

.004 files are short, and contain a standard header, followed by a data area containing not fully known information

Example Contents
07 [d3 ef 0a 4b] 17 87 08 00 35 [26 8e 0a 4b] 29 7b 08 00

The two 32bit values in brackets are obviously timestamps.

Information about this Document
This document was compiled and written by Mark Watkins <jedimark64@ (idontfeargoogle)>

This knowledge contained in this document represents the combined work of several hackers of the CPAPTalker's community.

Credit especially goes to Mike Hoolehan (see http://www.onkor.net) for the initial grokking into this weird and wonderful data format.

I expanded on this work by decyphering the ASV stuff and the remainder of the waveform headers.

Any terminology about Philips Respironics Trademarks belongs to them, (things like Flex types, machine and model names, etc..)
The author along with this work is in no way affiliated with Philips Respironics.. Just a hacker with OSA and an earnest desire to understand how CPAP machines (and the related health conditions) work.

This data format was lawfully reverse engineered by experienced hackers studying the raw SD card data very carefully. No guarantees are made regarding it's accuracy.
If you need to know any details with 100% certainty, I'm sure Philips Respironics will have a nice slimy NDA waiting for you.

Information deserves to be free! These machines are recording our breaths. I for one, reserve my right to use any such recordings made of my breath, along with any information derived from this, for what ever purpose I see fit..

This information is released under the Creative Commons.
Some more free document license stuff goes here..